Samson+ (Neptune)

Phase Change Thermal Forcing System

Operation Manual

Neptune

Revision: June 2025

Contents

1.	Introduction	2

1.1.	Standard Features	3

1.2.	Hazards	3

1.3.	Safety Features	3

1.4.	Facilities requirements	4

2.	General Dimensions	5

3.	Before you start	5

4.	Tool Operation	5

Basic operation	5

Clamp OPERTION

ClampAir  HOSE CONNECTION

CLAMP  UP/DN

5.	InTEC connection	8

6.	PID Parameters, Advanced Options	9

7.	Graph view	9

8.	Alarms	10

9.	Firmware	11

10.	Thermal Kit (Pedestal) Mounting	11

12.	Maintenance:	11

Appendix A	12

Appendix B – General Dimensions	15

Appendix C – Firmware Update	18

# Introduction

The new cooling/heating forcing system is targeted to answer customers’ demands to validate Si components in temperature ranges of -55ºC up to +150ºC in a cooling/heating envelope between 350W@-40ºC  to 1200W@0ºC  and 1300 watt @ 140c. The Samson+/Neptune thermal tool was designed to fit specific dimensions to fit into Intel Six Shot rack. ( 30x30x550cm). The tool thermal head is designed to fit  pneumatical force actuation up to 200 k”g and gimbaling with minimum KOV (Keep Out Volume).

Samson+  with Clamping device

- Samson+  is refered to 3 ph power supply.

- Neptune is refred to 230v 60hz power supply.

Samson with Clamp

## Standard Features

- Samson+/Neptune comes in three configurations:

- Neptune Water (PCWliquid cooling) 3liter@15c 230 volt 60 hz

- Neptune Water (PCW liquid cooling) 3liter@15c 230 volt 50 hz

- Samson+ Water (PCW liquid cooling) 3liter@15c 3 Phase 440v -480v 50/60hz  +Ground + natural. Phase to Phase is 480Volts  . The system is indefrence in the phase order  . The Delta -Y onnection is internal and has no relevance.

- Cooling power envelope of the tool is 350W@-40ºC(Tc) and 1200W@0ºC(Tc) with 10 mm pedestal Highet

- Temperature range -56ºC to +150ºC.

- Temperature change rate (hot/cold) is approximately 1.4 ºC/sec (without load).

- Temperature control ±0.2ºC from Set Point.

- The hose length between the thermal head and the control box is 230cm (220cm for the first models).

- Remote interface:option 1; using I2C comm’ protocol is specifically designed to support InTEC controller.option 2;TCP over UDP w/ python lib both can work with  Static-ip or DHCP.

- Defrost system activates below 20ºC.

- Dry Air purge is avilaible with Thermal head size 70x70m

## Hazards

- Each thermal tool is specifically designed to be powered by either of the following sources:

  - 230V 50Hz. 6A during operation.

  - 230V 60Hz max. 6A during operation.

  - 3 Phase 440 v and max 4A during operation.

- All control sub-systems and wiring use 24Volt DC.

- Internal refrigerant system contains pressurized refrigerant R452A.

- All pressurized system parts comply to commercial piping industry standards.

## Safety Features

- The operation of this device should be by authorized personnel only.

- The Samson+/Naptune tool intended be installed in a rack or a table.

- All system components comply to CE standards. See appendix A. MSDS for SUVA® R452A Refrigerant gas is used for cooling.

- https://www.fsw.uk.com/sites/www.fsw.uk.com/files/r452a_sds_0.pdf

- https://www.honeywell-refrigerants.com/europe/wp-content/uploads/2017/10/Solstice-452A-UK-english-1.pdf

- Maximum internal pressure is 15 Bar. Refrigerant weight is 300gr.

- The tool will power off when its thermal head temperature reaches 150ºC , or internal water leak detected

- For Samson+/Neptune In case of leak detection, the main water valve will close, and the tool will power off.

- Intel InTEC controller will turn the Neptune thermal tool off at 130ºC.

1.4 Safety Symbols

Below are examples of safety symbols and their meanings. They appear on certain tool parts.

- Caution – risk of electric shock

- Ground – connect power supply ground wire here

- Hot/ Cold – Tool thermal head surface can reach very cold and very hot temperatures. Be aware before touching the thermal head during operation.

! Temperatures can reach -40⁰C and +125⁰C

## Facilities requirements

- The Neptune should be powered through a circuit breaker used as the main overcurrent protection device and main disconnection device. A circuit breaker rated 230V, 16A or 120V, 20A ( be aware in rush current for 120v 60can be up to 50 amps ) So wires supply need to be

  - 230V 50Hz (inrush up to 30 Amp ) and max. 6A during operation.

  - 230V 60Hz (inrush up to 50Amp) and max. 6A during operation.

  - 3 phase 440 volt 40/60hz (inrush up to 25Amp) and max. 3.5A during operation.

    - System Type: 480/277v

    - Lile to line voltage :480v

    - Line to natural  Voltage: 277v

    - Note :Industrial lighting and Motors

    - Dry Air requirements  max 4  litter/min  due point -80c.

  - Water supply : 2.8 -3.2  lite/minute at 15c cooling capacity at least 1.5kw @ 15c

  - Regular Air Supply  6 bar 8 mm hose diameter to support thermal head pneumatics

.

- Neptune Water Chillr  tool requires a standard 208-230V 50 hz or 3Phase 480v 60 hz  outlet only, depending on a tool specification. PCW with a flow rate of 3.2±0.2 Litter/per Minute and a temperature of 15⁰C.

- Dry Air Flow recommended up to 2 liter per / min due point -80c

- Total cooling system efficiency is ~55%.

# General Dimensions

- See Appendix B

# Before you start

- Make sure that the tool thermal head doesn’t interfere with environment and the tool hose is  not bent with a minimum radius of 20cm (7.9 IN) to allow free flow.

The next 4 points points are relevant to intel equipment only.!!

- Verify that the InTEC cable (a gray cable) is securely connected to the tool control box

- Verify that the InTEC cable (a gray cable) is connected to InTEC – while InTEC itself is not switched on.

- Verify the TC (Thermocouple) connector is connected to its mating connector on InTEC

- Mount the tool UI LCD on any ferrous surface (Six Shot rack side wall or a cart side). For Neptune Water tool the UI LCD can be left attached to the tool control box.

- Verify the TC (Thermocouple) connector ,male is mounted to the blue famale conector along the main hose for T-case measurement and is well mounted to the pedestal

- Caution: Before connecting a power cord to the tool, verify the electricity voltage and frequency match both on a tool spec and a power outlet. The power requirements are written on a tool side identification panel.

Power outlet options:

Option 1: 230V- 50Hz, 8A MIN

Option 2: 120V 60 Hz, 16A MIN

Option 3: 3 phase 440/400 volt as  specified above.

- For Neptune Water tool, connect the PCW supply and return hoses.

PCW supply is marked with “S”. PCW return is marked with “R”.

Neptune Water tool Supply and Return hoses connection

If u used a chiller from Bonn Cooling be aware connect the Chiller hoses  the correct directions:

- Chiller OUT goes to Neptune S (Supply)

- Chiller IN goes to Neptune R ( Return to chiller)

# Tool Operation

## Basic operation

## CLAMP OPERTION

## Clamp Air coonection schematic

## Clamp Control panel

Up/Dn two nobes

Force ajusment nobe

## Samson + front panel

Main Fuse 20 A

Main switch

Main switch

Water

return

Water supply

Ethernet

Power Cord

- Plug in the  3ph connector as described above

- Switch the tool to “ON”  by moving the  green switch to the right green light shoud be show and the main LCD will be displayed.

- Tap “Main” in the upper right corner of the UI LCD.

Force ajusnent display

Set Temp

Main tab

Set

Temperature

- Make sure temperature Set Point is set to 25ºC, if not tap the temperature value next to the Set and change to 25ºC. Tap the Enter Key (left arrow) to enter the new temperature.

- Turn on the tool by tapping the  button on the UI LCD, after a few seconds (up to 30 seconds) the tool status will change to green indicating that it is ON.

- Wait at least 5 minutes until pressure and the four tips temperatures stabilize (25ºC).

- In order to view system parameters, enter the service screen by tapping Service and then System:

Tcase Reading

Service and System

buttons

Mode screen

In the Mode screen a user is able to change the defrost heating level at rate 0%, 25%, 50%, 75% and 100%. The default should be set to 75%.

There are other parameters in this screen, but their values should be changed only by a qualified maintenance personnel.

Caution: only a qualified maintenance personnel should modify these parameters as system operation and performance can be affected with an improper use.

Defrost Level:n

change to0%  25%-75%

Intec Screen

Intec screen will indicates Hot Gas con

If Hot gas will raise above 60c alarm will be triggered – meaning water supply is not sufficient  – should be in Temp range of 20-60c  and flow of  3 Liter per minute

# InTEC connection

- Verify the InTEC cable is connected both to the Neptune and to the InTEC.

- Turn the InTEC power switch on.

- After several seconds a message “Intec Control” will blink on the Neptune UI LCD indicating that InTEC controls the Neptune tool now.

Note: None of the Neptune tool parameters can be changed while InTEC controls the Neptune tool. In order to change the Neptune tool parameters, the tool should be changed to a LOCAL mode.

Blinking message indicating InTEC control

# PID Parameters, Advanced Options

The PID values can be modified from Service  Control screen

The PID values are to be used only by qualified maintenance personnel.

PID Parameters

Caution: only a qualified maintenance personnel should modify these parameters as system operation and performance can be affected with an improper use.

# Graph view

In order to see the Temperature Graph tap Graph on the main screen.

There are 2 parameters to view the graph:

- PV – Process Value

- SV – Set Value

PV- SV button displays the delta between process value and setting value.

PV button displays the actual temperature.

The Delta between process value and setting value

PV -Process Value

SV-  Set Value

PV

PV-SV

Scale lower limit

Scale upper limit

# Alarms

- Alarm screen logs alarm events during operation. All alarms are transmitted to the InTEC controller (if connected).

example alarms messages: I2C communication error, High temp, Low compressor pressure, Tip temp difference.

A critical alarm will shut down the machine: a thermal head temperature reaches 140ºC, a cooling fan is not functioning properly, water leak detected, a compressor temperature reaches above70ºC in Neptune Air, a condenser temperature reaches above 70 ºC (Neptune Water)

# Firmware

The Neptune tool comes with the latest available firmware straight from the factory. However, firmware updates are being released from time to time.

To update the tool firmware, connect an ethernet cable to the RG-45 connector which is located on a front panel of the Neptune control box.

For firmware update procedure see Appendix C

# Thermal Kit (Pedestal) Mounting

The thermal pedestal should be assembled following  provided procedures. Use gloves

Make sure that Samson surface is clean: push out Samson main piston like showing on the picture.



Verify that you receive correct pedestal and it compatible with you product. Verify that you have correct thermocouple. 
Make sure that thermocouple is attached well to pedestal.

Put a thermal grease (thin layer) on Samson surface. Use Thermal grease that compitable to your product. 



Attach pedestal and lock it with screws. After you press on a device at first time, tighten the screws again. And clead up any grease residure.



Connect the thermocouple to the Samson outlet. Make sure that cable passing smoothly and does not interfere with the movement of the main piston. Performe UP/DOWN tests and make sure that movement is smooth. Set purge pressure for your product request.

# Maintenance:

A Neptune tool should be Service for inspection and maintenance every 2 years. The inspection includes checkups for leak detection, valves operation and a compressor operation. For more information please contact Neptune tool -Intel product manager.

## Appendix A

## Appendix B – General Dimensions

Neptune Water control box.

A thermal head. This thermal head comes with both Neptune Air and Neptune Water tools.

## Appendix C – Firmware Update

- Tap Service  Info in lower right screen, the following screen appears:

DHCP

(highlighted)

IP Address

Firmware Version

MAC

Address

- Make sure the tool is in  mode.

- Download “LMFlash Programmer” software to a local PC from the following link (downloading firmware to Neptune can be carried out on a PC with DHCP/DIRECT link configuration).

LM Programmer Link:

http://www.ti.com/tool/LMFLASHPROGRAMMER

Must use build 1613

- Make sure that the needed firmware *.bin file is available on you PC.

- Disconnect Wi-Fi on your PC and change Internet Protocol Version 4 (TCP/IPv4) to use a specific IP (see the below picture).

- Connect LAN cable between the PC and the Neptune tool.

- Open the LM FLASH.exe

- Configure Quick Set to Manual configuration and then set Client IP address to 1.1.1.2 and Client MAC address to tactual Neptune MAC address (to see the MAC address navigate to Service --> Info)

- From the Program menu browse to the needed directory to select the FW file

- From the Neptune main screen change InTec to Local then touch turn off on the tool UI LCD to put a tool in a turned off mode.

- In UI LCD Navigate to Service --> Info to change the IP to Static IP and update IP to 1.1.1.2

- Go to LM Flash and press "program". It should take several seconds to have the new firmware "burned".

- On Neptune UI LCD there will "Firmware Update" message.

- The tool will restart. Go to Service --> Info and check the firmware version to display the installed version.